Use of the AMNIOQUICK test kits

| Document Author | Anne Penrice |
| --- | --- |
| Document Owner | Anne Penrice |
| Approved By | Clemora Thomas |
| Review Interval | 2 years |
| Location of Hardcopies | SH-DEL YH-G3 |

Contents:

1	Purpose and Principle	3

2	References	3

3	Equipment	3

4	Personnel Authorised to Perform Procedure	3

5	Sample Requirements (including COSHH Risk Assessment & First Aid)	3

6	Risk Assessment	5

7	Calibration	5

8	Quality Control	5

9	Method	5

10	Reporting of Results	5

11	Assay Performance & Known Limitations	6

# Purpose and Principle

The test kits are used to provide a rapid test for the detection of ruptured fetal membranes in pregnant women. Premature rupture of membrane ( PROM) is suspected in 5-10% of all pregnancies and may result in the preterm delivery and fetal infection. The leakage of amniotic fluid is not always detectable by clinical examination and may require a biological test. AMINOQUICK is one test that relies on the presence of the IGFBP-1 molecule.

Principle

A pair of IGFBP-1 monoclonal antibodies is used for IGFBP-1 detection. One is immobilized on the membrane of the test kit on the test line the other is attached to colloidal gold particles. During the sample migration IGFBP-1 in the sample forms an antigen-antibody complex with the colloidal gold antibody. This complex is the captured by the antibody on the test line and produces a purple line.

# References

- Package insert from AMNIOQUICK test kits

# Equipment

Test kit –product code 1090003

Kits should be stored at 2-30 and are stable until the expiry date shown. Once the test pouch is opened the test should be performed within an hour.

Material provided in the test kit

- Individual  wrapped test strip with desiccant

- Sterile Decron Vaginal swab

- Unit dose vial containing diluent

- Dose vial rack

Also required

- Timer

Control solutions –product code 6090001

Kits and control solutions are ordered from

BioSynex

# Personnel Authorised to Perform Procedure

Any member of staff who has completed the AMNIOQUICK face to face training/Quiz

# Sample Requirements (including COSHH Risk Assessment & First Aid)

Use the dacron swab to collect secretions from the vaginal wall. Open the swab bag and place the swab into the vagina to a depth of 5cm and leave for 1 minute. If a speculum is used the secretions may be collected by leaving the swab in contact with the vaginal wall at the level of the posterior formix for 15 seconds.

The swab must be placed in the dilution tube immediately after collection. Once the swab is in the extraction tube it may be stored for 6 hours at room temperature.

- All human bodily fluids samples must be treated as potentially BIO-HAZARDOUS.

- Approved Personal Protective Equipment (PPE) including gloves and eye-protection must be worn when handling open

- samples or derivatives thereof.

|  | When performed according to the protocol detailed in this SOP, and in conjunction with adherence to Trust Policies and Good Laboratory Practice, the handling of patient samples represents minimal risk to staff. |
| --- | --- |

Exposure to Bio-Hazardous Material

In the event of a needle stick injury or accidental blood splashes to eyes or mouth:

- If skin has been punctured encourage bleeding by gently squeezing.  Wash with soap and running warm water then dry and dress the wound.

- Splashes to the eyes: irrigate eyes thoroughly with eye wash / saline

- Splashes to the mouth: gargle with drinking water (avoid swallowing)

Contact the Occupational Health Department / Emergency Department for guidance and report all adverse incidents to your line manager / complete a DATIX form.

Disposal of Patient Samples

Ensure compliance with the trust policy

The following first aid guidelines may be applied to all the substances detailed in this SOP.

Eyes: Irrigate thoroughly with water. At least 10 minutes is the recommended duration. Sterile saline is also available at the eye wash stations.

Lungs: Remove from exposure, rest and keep warm.

Skin: Wash substance off skin thoroughly with water. Remove contaminated clothing and wash before re-use.

Mouth: Wash out mouth thoroughly with water and give plenty of water to drink.

Remember – If at all concerned about the nature or severity of the problem, SEEK MEDICAL ADVICE.

# Risk Assessment

# Calibration

None required

# Quality Control

A negative and positive control sample should be run on each box of test strips to check the strips are working correctly

# Method

- Bring the complete kit to room temperature prior to testing

- Open the vial containing the diluent and place it in the rack tube to keep it vertical

- Dip the swab into the diluent and rotate for 10 seconds. The swab may then be  discarded

- Place the test strip arrows pointing downwards into the vial making sure it rests in the bottom of the tube. The strip must be left in the tube for 10minutes.

- After 10 minutes read the test results from the strip. Results should not be read after 15 minutes and the vial and strip should be discarded in compliance with local policies.

# Reporting of Results

The test strips should be read in good light.

Positive tests for amniotic fluid requires

- A purple line in the control line area of the test strip

- Any intensity purple line in the test line area of the test strip.

Negative Test for amniotic fluid

- A purple line in the control line area of the test strip.

No line should be visible in the test area of the test strip

Invalid test

- Absence of a purple line in the control line area of the test strip i.e no lines visible on the test strip or a line visible only in the test area of the test strip.

# Assay Performance & Known Limitations

Detection limit

The detection limit for the AmnioQuick test kit is 5ng/ml

False positive results

May be caused by

- IGFBP-1 in bleeding

- Decidual cells on the when the cervix is mature enough

False negative results

- Test carried out >12 hours post amniotic leak

Interference studies

The following do not interfere with the AmnioQuick test

- Capillary blood

- Urine

- Semen

# Text boxes

Changes from last version of this document

First entry in Q-Pulse

COSHH and Risk Assessment

When performed according to the protocol detailed in this SOP, in conjunction with the safety precautions outlined and good laboratory practice, this method presents minimal risk to staff and patients.